## GeneralNotes
CPC SN: | 3789241001
EC SN: | 3082002408002
DMA SN: | 30832411001
Observations:
All times/data are reported in PST

## MonthlyQCs
Date | Start Time (PST) | End Time (PST) | Roof flow (SLPM) | EC flow (SLPM) | EC flow (indicated) | Sheath Flow (SLPM) | CPC Flow (SLPM) | Dryer gauge (inHg) | Zero air Test | Notes
2025-04-01 | 12:00 |  | 16.02 | [6 empty] | Inlet flow at install, logged as LPM.
2025-04-17 | 8:43 | [7 empty] | Done | Flow and zero checks performed; values and result not logged.
2025-08-13 | 10:20 |  | 16.4 | 1.42 | 1.48
2025-12-22 | [8 empty] | Done | Flow and zero checks performed; values and result not logged.
2026-01-08 | [9 empty] | Leak and flow checks passed after RH sensor install; values not logged.
2026-01-22 | [9 empty] | Monthly QC logged with no values or result.
2026-02-19 |  |  | 16.9 | 1.39 | 1.41 |  |  |  | Pass
2026-03-19 |  |  |  | 1.42 |  |  | 1.44 |  | Pass
2026-03-30 |  |  | 16.8 | 1.422 |  | 11.88 | 1.466 |  | Pass | Pre-shipment for annual maintenance. Sheath 11.85-11.91 SLPM on both - and + sides (Alicat); midpoint shown.
2026-05-01 |  |  | 16.1 | 1.42 | 1.42 |  |  |  | Pass | Opening QC after service. Zero not true zero (~5-10 particles, 7-30 nm).
2026-05-14 |  |  | 17.1 | 1.41 |  | 11.7 |  |  | Pass
2026-06-11 |  |  | 16.9 | 1.44 |  | 11.7 |  |  | Pass
2026-07-09 |  |  | 17 | 1.42 |  | 11.7 |  |  | Pass
2026-08-06 |  |  | 16.3 | 1.417 |  | 11.6 | 1.456 |  | Pass | Roof flow 17.0 before changing critical orifice black 3 LPM -> white 2 LPM; 16.3 after.
2026-09-01 |  |  | 16.9 | 1.409 |  | 11.56 | 1.449 |  | Pass | Roof flow 16.1 before changing critical orifice white 2 LPM -> black 3 LPM; 16.9 after.
2026-10-02 | 11:08 | 11:41 | 17.02 | 1.429 | 1.43 | 11.62 | 1.464 | -18 | Pass

## EventsOrVisits
Date | Time (PST) | Cleaned Impactors | Notes
2025-03-31 | 14:00 |  | Initial install.
2025-04-01 | 12:00 |  | Sampling inlet operational - 16.02 LPM.
2025-04-02 | 15:00 |  | Changed logging from 64 to 128 channels/decade to match 10WEL SMPS.
2025-04-11 |  |  | (MCH) Classifier would not start from AIM (greyed-out restart arrows; message that DMA voltage must be between 0 and 10000 V). CM logged in remotely: sheath flow in setup file needed to be 12 LPM instead of 15 LPM. Start-up procedure: in AIM, connect to the SMPS s/n, load the SeaTac airport setup file, change sheath flow from 15 to 12, then start the device. Scanning now.
2025-04-17 | 8:43 |  | Cleaned impactor and PM10 & PM2.5 cyclones. Flow and zero checks performed (values not logged).
2025-05-05 | 6:30 |  | Restarted AIM software - white window, not collecting data.
2025-05-29 | 13:30 |  | Setting change - AIM DMA-to-CPC tube length was set at 33.02 cm but was measured at 38 cm. Made adjustment in settings and impact on 3/31/25-5/29/25 data is expected to be limited
2025-06-16 | 14:00 |  | SMPS sent to TSI for yearly cleaning/calibration, no 18THA data until 2025-07-11
2025-07-11 | 10:00 |  | CPC back from TSI with flow mode switched to low flow (0.6 LPM) while AIM was set at 1.5 LPM. Invalidated data for that day between 10:00 and 11:00 AM. Fixed flow back to 1.5LPM in CPC.
2025-07-17 | 8:20 |  | CPC switched again to 0.6 flow during a running AIM session. AIM flow remained at 1.5 LPM resulting in concentrations biased low and higher particle sizes. Fixed on 2025/7/21 8:15. Data invalidated.
2025-08-13 | 10:20 |  | (CM) Cleaned impactor and PM10 & PM2.5 heads. Flows checked (see MonthlyQCs).
2025-08-13 | 12:10 |  | CPC rebooted and switched back to 0.6 LPM flow at reboot and came unnoticed. Restored only to correct 1.5LPM on 2025-08-27 10:20. Data invalidated.
2025-09-30 | 6:15 |  | Restarted AIM software - not working (white screen).
2025-10-13 | 6:15 |  | Restarted AIM software - not working (white screen).
2025-10-16 | 10:00 |  | Restarted AIM software - looks like a Windows update; software did not start after the laptop likely restarted.
2025-10-28 |  |  | Power outage on 10/26.
2025-12-10 | 7:20 |  | Restarted AIM software - frozen (white screen).
2025-12-22 |  |  | Cleaned impactor and PM10 & PM2.5 cyclones. Flow and zero checks performed (values not logged).
2025-12-26 |  |  | SMPS stopped 12/24/25 23:41 - AIM screen completely blank. Restarted AIM; SMPS working again.
2026-01-08 |  |  | Installed RH sensor at the EC inlet. Leak and flow checks passed.
2026-01-22 |  |  | Monthly QC (no details logged).
2026-02-05 |  |  | Cleaned impactor and serviced water reservoirs.
2026-02-09 |  |  | Windows crashed - a reboot restored service.
2026-02-13 | 7:40 |  | Noticed SMPS stopped 2/12 21:00 - AIM software frozen. Restarted it this morning.
2026-02-19 |  |  | Monthly QC passed. Cleaned impactor and cyclones. Windows updated and paused again.
2026-02-27 | 14:15 |  | Installed Partector 2 Pro (P2P) and connected to probe. Changed 3 LPM orifice to 2 LPM, but P2P only draws 0.5 LPM, so overall flow through the cyclones will be less than before. Started new file on the SMPS.
2026-03-11 |  |  | AIM Java error - restarted AIM.
2026-03-19 |  |  | Monthly QC passed. Cleaned impactor and cyclones.
2026-03-30 |  |  | Restarted AIM and Partector data collection - Windows security update on 3/29 prompted a laptop reboot.
2026-03-30 |  |  | Taking down SMPS for annual maintenance. Pre-shipment QC passed (see MonthlyQCs); sheath flow measured with Alicat since TSI 4140 flow meter barbs are too small for sheath flow tubing. Purge mode active for 10 min, then turned off and ran for an additional hour.
2026-05-01 |  |  | Returned from service. Opening QC: total flow 16.1 SLPM, EC flow 1.42 (indicated 1.42). Zero check OK but not true zero: ~5-10 particles in the 7-30 nm range, varying in location over 8 scans.
2026-05-05 |  |  | Changed scan settings to 46 s scan + 12 s purge, 64 channels/decade.
2026-05-07 | 13:00 |  | Switched from DI to distilled water.
2026-05-08 |  |  | (CM) AIM AutoExport settings: added missing upper bin at 392 nm (385 nm for 10WEL) and changed export resolution from 128 to 64 channels/decade.
2026-05-11 |  |  | Installed Pratector2Pro (P2P) instrument and adjusted orifice to accommodate. Using it to confirm suspected SMPS sizing problem (compressed/wrong-size bins).
2026-05-12 |  |  | P2P geometric mean is ~2x the SMPS, confirming sizing problem. Reboot of laptop, EC, and CPC seems to have fixed it; cause unknown, started around May 5 (or possibly earlier).
2026-05-14 |  |  | Monthly QC passed
2026-05-29 |  |  | Emptied drain bottle.
2026-06-11 |  |  | Monthly QC passed.
2026-07-09 |  |  | Logger blue-screened and rebooted the previous night. Monthly QC passed.
2026-08-03 | 6:45 |  | AIM software froze with error message; rebooted fine after restarting software. Same error as at 10WEL that morning - not sure if related.
2026-08-06 |  |  | Monthly QC passed. Changed critical orifice from black 3 LPM to white 2 LPM (probe inlet 17.0 -> 16.3 SLPM). Windows updated and entire system rebooted.
2026-08-10 | 6:15 |  | Restarted AIM software (stopped on 8/9), likely due to Windows update. Paused Windows updates until 8/31/26.
2026-09-01 |  |  | Monthly QC passed. Changed critical orifice from white 2 LPM back to black 3 LPM (probe inlet 16.1 -> 16.9 SLPM). Windows updated.
2026-09-18 |  |  | Windows updated; update KB5129195 prevented AIM from starting, so uninstalled it. Manually locked Peplink modem to band 13 (also tried bands 4 and 66 as recommended; 13 gave best results).
2026-10-02 | 11:08 |  | Backed up DB
2026-10-06 | 10:45 |  | Remotely started AIM - stopped overnight likely due to a laptop reboot.